## Slide 1
Alternative communication with single-muscle EMG
VE458
Biomedical Instrumentation and Design
Final Project Proposal
Group 7
Lu Xiuneng Lin Junwei
Wang Yuxuan Tommaso Cemmi
2017/11/27

## Slide 2
Index
01. Diagram of Our System
02. Components & Costs
03. Hardware Part
04. Software Part
05. Demo & Results
06. Suggestions & Conclusions
07. Q & A

## Slide 3
Lu xiuneng
Amyotrophic Lateral Sclerosis (ALS)
A disease that causes the death of neurons which control voluntary muscles
01
04
02
03
Stiff Muscles
Muscles
Decreasing in
Size
Difficulty in
Speaking,
Swallowing
Difficulty in
Breathing
30%

## Slide 4
Lu xiuneng
Aim
To help those without limbs, cannot speak but at least one of his muscle are active to communicate with others.

## Slide 5
01
Lu Xiuneng
Diagram of Our System

## Slide 6
Lu Xiuneng
Design - Overview
http://vibrationtrainingdevice.com/VG-Evolution/emg-elettrodi-b.jpg
http://sine.ni.com/images/products/us/04231404_m.jpg
https://s3.amazonaws.com/siteninja/multitenant/images/6912/images/slide/Screen-hand-540x300_1_.jpg_1266619279?1340887223

## Slide 7
Lu Xiuneng
A Real Image

## Slide 8
02
Lu Xiuneng
Components & Costs

## Slide 9
Lu Xiuneng
Components & Costs
Components | Quantity | Cost per unit (CNY) | Cost (CNY)
AD620 Amp | 1 | 7.8 | 7.80
LM741 Amp | 1 | 2.5 | 2.50
Wires | 1.5m | 1.18/m | 1.77
Capacitors | 2 | 0.75 | 1.50
Resistors | 4 | 0.44 | 1.76
DAQ board | 1 | 3048 | School Provided
Total |  |  | 15.33

## Slide 10
03
Lu Xiuneng
Hardware Part

## Slide 11
Functions & Schemetics
Functions: Get the EMG signal, amplify it and filter the noise.
Expected gain: ~300
Expected Cutoff frequency: 50 Hz & 450 Hz
Schemetics:
Lu Xiuneng

## Slide 12
Lu Xiuneng
Test Results for Hardware
We let Vin = 50mVpp, and here is the results:
Frequency (Hz) Vout (V) Gain
10 3.50 70.0
20 6.51 130.2
30 8.9 178
40 10.7 214
50 12.1 242
100 14.5 290
120 14.6 292
150 14.6 292
200 14.2 284
300 13.1 262
450 11.3 226
600 9.4 188
1000 6.2 124
2000 3.4 68
Actual Gain: 290; Cutoff frequency: 35Hz and 480Hz

## Slide 13
04
Lin Junwei & Wang Yuxuan
Software Part

## Slide 14
Lin Junwei
Functions of Software Part
Encode instructions
  Long signal represent 1, short represent 0
Execute instructions
  Different modes & instructions
Display the results

## Slide 15
Lin Junwei
Common Word list
A wordlist with 4300+ most common used words
When first several alphabet are input, the suggest word will appear

## Slide 16
Lin Junwei
LabVIEW Block diagram (II)
EMG Signal Acquiring & Decoding

## Slide 17
Wang Yuxuan
Finite State Machine

## Slide 18
Wang Yuxuan
LabVIEW Block diagram (III) -- FSM
LabView FSM Stages
Input Signal
Count/Reset Digit
Add digit 1/0
Select/Delete/Enter/Pick word
Check Code/Count digit for word
Enter sentence 1/2
Clear All/Sentence
Check Code
Exit

## Slide 19
Wang Yuxuan
LabVIEW Front Panel
17

## Slide 20
Wang Yuxuan
Input Signal

## Slide 21
Wang Yuxuan
Count Digit

## Slide 22
Wang Yuxuan
Check Code

## Slide 23
Wang Yuxuan
Check Code for Word

## Slide 24
Wang Yuxuan
Word filter

## Slide 25
06
Tommaso Cemmi
Demo & Results

## Slide 26
Tommaso Cemmi
Expected Result
  Show sentence on the
  screen
  Communicate any
  message
No limited list of words
  Optimize input method
  Suggesting words that
  can be selected
  Selecting common
  sentences from a
  list of options
  Able to input letters
  Able to delete wrong
  inputs

## Slide 27
Tommaso Cemmi
A Huge Advantage
Hawking's wheel
Our design
Costs: $950,00 - $ 1,000,000
Costs: ~$2.4
http://baijiahao.baidu.com/s?id=1563218123737396&wfr=spider&for=pc

## Slide 28
06
Tommaso Cemmi
Suggestions & Conclusions

## Slide 29
Tommaso Cemmi
Further Improvement
Use more sensitive filters
Enlarge the dictionary
Optimize the searching algorithm
Add annother mode to enter sentence by the first alphabet
of each word

## Slide 30
Tommaso Cemmi
Conclusions
  LabVIEW Base System
  01
  VALUE
  One input
  channel
  02
  VALUE
  Able to fast
  output common
  sentence
  03
  VALUE
  Support input
  complicated
  word or sentence
  04
  VALUE
  Support input
  complicated
  word or sentence
  04
  VALUE
  Word suggestion system
  05
  VALUE

## Slide 31
Tommaso Cemmi
Teamwork tasks
Lu Xiuneng
1
Wang Yuxuan &
Lin Junwei
2
Tommaso Cemmi
3
Hardware design:
Software design and FSM:
Systems and DAQ design:

## Slide 32
Q & A

## Slide 33
Questions & Answers

## Slide 34
Thank you